Evidence-Informed Design of Total Wrist Arthroplasty: Clinical Stratification, Mechanical Modelling, and Validation Priorities

From patient context to model conditions and postoperative assessment

Abstract

Total wrist arthroplasty (TWA) preserves useful motion while restoring stability and durable fixation across patients with different bone stock, deformity, and soft-tissue capacity. This critical narrative review organizes clinical and biomechanical evidence around three design decisions: defining the reconstruction problem for each patient, specifying model conditions for comparing candidate designs, and selecting postoperative measurements that can test predicted responses. PubMed and Europe PMC searches through September 2026, supplemented by citation chaining, informed a synthesis spanning clinical requirements, patient characterization, computational mechanics, design, manufacturing, and assessment. We provide a patient-stratification matrix, a model-condition reporting template, and an evidence-to-decision map. Wrist studies show that load transfer, carpal motion, and fixation response vary with task, implant geometry, and retained tissues. Patient-specific imaging and simulation support conditional comparisons of candidate designs, and clinical reports describe customized implants in selected cases. Wrist-specific CT-based migration studies quantify component motion; ultrasound interface methods remain at the experimental stage. The central research priority is to connect reproducible loading and model assumptions with independent mechanical measurements and patient outcomes. This synthesis makes those inputs, decisions, and evidence gaps explicit.

Keywords: total wrist arthroplasty; implant design; carpal kinematics; medical imaging; multibody simulation; finite element analysis; patient-specific implants; design validation

1 Introduction

1.1 Clinical context and design challenge

End-stage wrist arthritis requires reconstruction of a compact joint system in which useful hand function depends on motion, stability, and load transfer [1, 2]. Carpal motion emerges from interacting articular geometry, ligaments, and musculotendinous forces. Its function emerges from coupled rotations and translations, migrating contact and position-dependent load sharing [3–5]. Total wrist arthroplasty (TWA) changes this architecture and must establish a mechanical environment in which the retained tissues and implanted components work together [6, 7]. The relationship between anatomy, bearing geometry, soft-tissue balance and fixation therefore defines the reconstruction problem.

Successive implant generations have improved pain relief and preserved motion, while revision, loosening, instability, fixation failure, and polyethylene wear remain reported outcomes [8–11]. Failure patterns vary across implant generations and architectures. Constrained designs transfer greater demands to fixation, whereas less constrained designs depend more on retained soft-tissue stability [1, 6]. Contemporary systems also report carpal-side loosening and late wear-related failure [10, 12]. These observations place fixation, bearing behavior, bone preservation, and soft-tissue balance within the same design decision.

1.2 Review rationale scope and literature synthesis

Mechanical measurements explain why acceptable component placement can coexist with altered function. Implanted wrists can exhibit changes in center-of-rotation (COR) trajectory, carpal height, contact migration and tendon moment arms [13–15]. These changes influence contact pressure, impingement, fixation moments and interface motion [16, 17]. Clinical outcomes also depend on disease, bone quality, activity, surgical execution and experience [18, 19]. Design evaluation must connect the mechanical response to this clinical context, so that a proposed improvement has a defined functional purpose and a measurable consequence.

Building that connection requires anatomical measurements, loading estimates and outcome observations to inform the same design decisions. Imaging, motion tracking, multibody simulation (MBS), finite element analysis (FEA), and construct testing provide complementary evidence about anatomy, loading, contact, and fixation [20–23]. The evidence becomes useful for design when these methods share defined tasks, coordinate systems, and assessment quantities. A measured trajectory can then constrain a model or geometry, and a predicted mechanical benefit becomes clinically informative when it can be validated by the independently measured response. More specifically, wrist motion is quantified using dynamic imaging and biplanar videoradiography, while segmentation and statistical shape modelling support anatomical reconstruction [24–27]. MBS and FEA estimate muscle forces, contact mechanics, and fixation loading under specified conditions [3, 16, 17, 22]. Customized implants demonstrate design and manufacturing feasibility in selected wrists [28, 29]. The remaining task is to connect these capabilities through patient-specific requirements and measurements that can test predicted performance.

This review contributes three linked tools: 1) a patient-stratification matrix that connects clinical findings to design and model choices; 2) a model-condition template for interpreting MBS–FEA comparisons; and 3) an evidence-to-decision map that distinguishes measured references, predicted responses, and acceptance criteria. The six-stage synthesis (Fig. 1) organizes these tools across clinical requirements, anatomical and kinematic characterization, computational analysis, implant design and optimization, manufacturing, and assessment. Each stage produces information that supports downstream decisions, while experimental and postoperative observations refine upstream assumptions. The central question is which clinical observations and mechanical predictions can guide TWA design, under what conditions, and what evidence can test their clinical meaning.

The critical narrative synthesis draws on PubMed and Europe PMC searches through September 2026, supplemented by citation chaining. Search terms paired total wrist arthroplasty, total wrist replacement, wrist prosthesis, and wrist implant with imaging, modelling, customization, manufacturing, and surveillance concepts. Studies reporting clinical outcomes or validated measurements informed the evidence map; studies from other joints or implant systems were retained as indirect methods evidence and identified as such. Selection and stage assignment were interpretive, and no prospective protocol or study-level screening log was created. Quantitative results are reported with their study population, device, task, or model conditions where these were available. The synthesis evaluates how the evidence supports design decisions and where the relevant measurement or clinical link remains to be tested [1, 2, 8].

Tab. 1 presents selected clinical outcomes and biomechanical measurements as separate evidence domains. Each value is interpreted within its study population, implant, and task.

Tab. 1. Selected clinical outcomes and wrist mechanical measurements

| Evidence domain | Study/population | Reported outcome | Use in this review | Interpretive scope |
| --- | --- | --- | --- | --- |
| Clinical survival | Krukhaug et al. {Krukhaug 2011}; registry cohort spanning implant generations | 5-year survival 78% (95% CI 70–85); 10-year survival 71% (59–80) | Historical TWA context | Interpret within registry era and cohort definition |
| Clinical survival | Newton et al. {Newton 2023}; Universal 2 long-term series | 10-year survival 78%; 22% revised; late wear-related loosening proposed | Device-specific follow-up | Long-term result for this series; mechanism is proposed |
| In-vivo mechanics | Akhbari et al. {Akhbari 2020}; TWA and healthy wrists under matched tasks | COR excursion approximately 17.1 mm after TWA and 7.2 mm in healthy wrists | Matched task and algorithm comparison | Descriptive motion reference; no outcome-linked target |

Fig. 1. Proposed six-stage synthesis linking patient requirements, characterization, mechanical analysis, design, manufacturing, and assessment. Arrows show information transfers and feedback proposed by this review; the evidence supporting each transition is summarized in Table 6.

2 Clinical and biomechanical requirements

Patient selection defines the feasible reconstruction and the measurements needed to assess it. Disease pattern, available bone, retained soft tissues, DRUJ status, activity, and revision context shape both the design space and the loading assumptions. Table 2 maps requirements to measures; Table 3 links clinical strata to design and model decisions.

Tab. 2.  Requirements, quantities, and evidence status

| Requirement | Measure | Evidence basis | Role in design | Status and interpretation |
| --- | --- | --- | --- | --- |
| R1 Functional motion | Prioritized task trajectory; Ryu daily-activity values are a healthy reference {Ryu 1991} | Healthy activity data and TWA reviews {Reigstad 2025; Clementson 2024} | Define comparison tasks and patient goals | Reference values; no universal postoperative pass value |
| R2 Load transfer | Task-specific contact and force path; native 80/20 is posture-dependent {Eschweiler 2022} | Native wrist studies and TWA FEA {Gislason 2017; McCullough 2012} | Compare load paths under matched tasks | Descriptive reference; reconstructed-wrist limits require evidence |
| R3 COR behavior | Task-based trajectory and excursion {Akhbari 2020} | In-vivo TWA and healthy-wrist comparison | Evaluate motion and placement sensitivity | Observed difference; no validated clinical threshold |
| R4 Fixation | Normal/tangential interface motion at stated sites and loads | Porous implant studies and wrist FEA {Pilliar 1986; Bobyn 1999; Gislason 2017} | Compare defined interface states | Candidate ranges are condition-specific, not universal TWA criteria |
| R5 Soft-tissue compatibility | Tendon clearance, carpal height, moment arms, retained tissues | Cadaveric and simulation studies {Bain 2025; Glanville 2025; Hooke 2015} | Include tissue constraints in geometry and tasks | Mechanistic evidence; patient-level benefit requires assessment |
| R6 Bone and revision | Resection, fixation corridors, extraction, and salvage options | Salvage literature {Cash 2025; Berber 2020} | Preserve feasible future procedures | Descriptive design constraint; no common numeric criterion |
| R7 Surveillance | Precision and repeatability of alignment/migration measurement | Wrist CT and indirect RSA/MAR methods {Lind 2026; Koo 2026; Jürgens-Lahnstein 2026} | Select measurement protocol and implant features | Measurement precision is distinct from diagnosis or prognosis |
| R8 Implementation | Dimensions, material, surface, and process conformance | Manufacturing studies and custom TWA reports {Mani 2024; Wei 2025; Cai 2023} | Verify that the manufactured device matches design intent | Qualification is device- and process-specific |
| R9 Clinical outcome | Pain, grip, motion, PROMs, complications, revision, and survival | Validated instruments and cohorts {MacDermid 1998; Chung 1998; Pappalardo 2026} | Test patient benefit and durability | Interpret by indication, comparator, definition, and follow-up |

2.1 Native wrist mechanics as a functional reference

The wrist comprises the distal radius and ulna, eight carpal bones and the proximal metacarpals. Interacting articular geometry and ligaments coordinate motion across these structures [5, 30]. In the neutral native wrist, approximately 80% of axial load passes through the radiocarpal region and 20% through the ulnocarpal region; the distribution varies with posture and measurement conditions [3]. Contact locations migrate during wrist motion [5]. Dart-throwing motion couples radial extension with ulnar flexion and commonly involves substantial midcarpal motion with reduced proximal-row rotation [5, 24]. These features establish the need to assess coupled motion and changing load paths.

Clinical range-of-motion measurements summarize hand motion relative to the forearm, while each carpal bone follows a six-degree-of-freedom trajectory [4, 5, 23]. A classic analysis of daily activities identified a functional envelope of approximately 60° extension, 60° flexion, 40° ulnar deviation and 20° radial deviation [3, 23]. These values provide task-based reference measurements. Reconstruction targets should reflect the activities prioritized by the patient and the motion achievable within the retained anatomy.

2.2 Musculotendinous and soft tissue constraints

Wrist flexors and extensors act through position-dependent moment arms and co-contract to stabilize the carpus during grip [3, 31–33]. Carpal height and COR position influence tendon excursion, tension and mechanical advantage [15, 16, 32]. Bearing geometry and component placement must therefore be evaluated together with the soft tissues that supply motion and stability.

Local tendon clearance also deserves assessment. A five-case surgical series of distal intersection tenosynovitis described tendon lesions associated with underlying wrist joint or bone abnormalities [34]. This adjacent wrist evidence supports examining tendon paths and potential contact with the reconstructed anatomy when defining soft-tissue clearance requirements.

2.3 Disease and patient specific constraints

Disease pattern defines the anatomy and tissue capacity available for reconstruction. Inflammatory arthritis may combine ligament attrition, carpal collapse, and osteopenia; osteoarthritis and post-traumatic destruction produce different patterns of joint damage and asymmetry [1, 2, 8]. Candidacy discussions therefore consider deformity, extensor function, soft-tissue integrity, infection history, bone stock, activity, and prior procedures[35]. Arthrodesis remains an important clinical comparator when the value of preserved motion is weighed against stability and durability [36]. These factors determine which reconstruction options and design constraints should enter the analysis.

The ulnar column contributes to load transfer, while the Distal Radio-Ulnar Joint (DRUJ) enables forearm rotation [37, 38]. DRUJ retention or reconstruction changes the anatomical and mechanical conditions represented in the plan. The operative strategy and model should therefore describe the same ulnar-column structures and forearm tasks.

2.4 Translating clinical needs into engineering requirements

Tab. 3. Patient strata and their implications for TWA design

| Clinical stratum | Findings to document | Design and model implications | Decision question |
| --- | --- | --- | --- |
| Inflammatory arthritis | Bone stock, carpal collapse, ligament and tendon status | Available fixation, retained constraints, tissue clearance, and task assumptions[1] | Which structures and functions can the reconstruction rely on? |
| Post-traumatic destruction | Asymmetric bone loss, prior procedures, alignment, and soft-tissue injury | Reconstruction target, fixation corridors, resection, and revision options[28] | Which defect prevents a standard component from meeting the stated requirement? |
| Degenerative arthritis | Joint damage, bone support, activity, and symptoms | Task selection and comparison with available clinical alternatives[36] | Which patient-prioritized task defines useful motion? |
| Limited carpal-side support | Cortical support, defect location, and previous fixation | Component footprint, screw paths, interface assumptions, extraction, and salvage[39, 40] | Can the proposed fixation be achieved while retaining a revision path? |
| DRUJ or soft-tissue involvement | DRUJ procedure, forearm rotation, tendon and ligament function | Ulnar-column boundary conditions, retained stability, and tendon paths[37, 38] | Which tissues and articulations are included in the reconstruction? |

The matrix structures clinical questions addressed in the cited literature. Each option is interpreted within the patient's anatomy, operative plan, and available clinical evidence.

Clinical needs become engineering requirements through measurable quantities. Functional motion is defined by prioritized tasks and coupled trajectories [24, 41]; stability by bearing constraint in relation to retained tissues [17, 32]; load transfer by contact and force paths through the radial and ulnar columns [30]; and fixation by interface motion under a stated load [16]. Bone preservation, extraction features, and revision options define additional design constraints [39, 40]. Each requirement needs a quantity and assessment method whose evidence supports the intended decision.

3 Anatomical reconstruction and quantitative motion characterization

Characterization of the native wrist supplies the geometry, tissue information, and motion measurements required by the clinical question. Table 6 maps representative evidence to these uses and identifies methods transferred from other settings.

3.1 Imaging for anatomy and motion

Imaging selection follows the quantity required by the design decision [4, 25]. CT characterizes bone morphology and, with appropriate acquisition and calibration, supports density-based material assignment. Dynamic four-dimensional CT and biplanar videoradiography quantify time-resolved bone motion; MRI depicts cartilage and soft tissues subject to motion and metal artefact [42–44]. Ultrasound scores characterize tenosynovitis in inflammatory disease and support soft-tissue assessment [45]. They do not directly provide tendon or ligament mechanical properties. Each protocol should be selected for its spatial and temporal resolution, field of view, radiation exposure, and artefact relative to the target quantity.

Dynamic acquisition should reproduce a defined task and provide poses or trajectories matched to the model input [23, 42, 46]. Tracking pipelines register sequential images to estimate bone pose. Their accuracy and repeatability apply to the tested acquisition protocol, anatomy, and motion range [26].

3.2 Anatomical reconstruction and model generation

Anatomical reconstruction proceeds through resampling, artefact correction, segmentation, and registration while retaining source images and processing versions [25]. Level set, shape-and-pose models, and deep learning provide complementary segmentation strategies [27, 42, 43, 47–50]. However, segmented accuracy is limited to internal limitation of imaging methodology, e.g., erosion, fusion, partial-volume effects, and metal artefact can change fixation surfaces and landmarks [42, 47, 48].

Segmentation quality is assessed using quantities tied to the intended decision. Local surface error informs fit and resection planning; landmark and coordinate-system repeatability inform alignment and motion analysis [51]. Articular boundaries, cortical support, and fixation corridors require particular attention because local geometric error can alter the design.

Segmented volumes become surface or volumetric models through geometry-preserving smoothing, component repair, and mesh-quality assessment [51]. Surface models support coverage and collision analysis; volumetric meshes support material assignment and FEA; anatomical landmarks support kinematic analysis; and watertight geometry supports manufacture [22, 52]. Cortical and articular boundaries should be represented according to the analysis, with cartilage-thickness uncertainty carried into contact assessment [53, 54].

Reconstruction of deficient anatomy begins with a specified anatomical target [27, 28, 47]. Contralateral imaging provides an individual reference when the opposite wrist is suitable; statistical shape models estimate population-informed geometry from partial observations; surgeon-guided correction records planned alignment and operative constraints [55–57]. The source population and reconstruction uncertainty accompany the resulting model. CT-based bone-property assignment requires scanner-specific calibration from attenuation to density and elastic modulus [16, 22].

3.3 Motion characterization and design inputs

Quantitative motion analysis requires reproducible coordinate systems across participants, sides, and sessions [4, 46]. Bone-centroid trajectories describe translation, while Cardan angles or quaternions describe orientation. Helical-axis parameters characterize rotational axes over defined motion intervals [5, 23, 58]. Reports should specify the reference frame, rotation order, sign convention, motion interval, filtering, and tracking error for each measure.

A bone's geometric centroid provides a translation descriptor, while its COR describes rotational behavior; both depend on the reference frame and measurement method [4, 55, 58]. External contours inform component footprint, resection coverage, stem entry, edge clearance, and soft-tissue space [54, 56]. Reporting centroid and orientation together links motion to the geometry that constrains a candidate design.

The characterization package contains reconstructed surfaces, coordinate systems, calibrated bone-property maps, and task-based motion measurements [52, 59]. Centroid trajectories, orientations, and rotational descriptors serve as prescribed inputs or comparison targets according to the model question [22, 58, 59]. Segmentation, registration, and tracking errors define the uncertainty carried into subsequent analysis.

Imaging and tracking studies report protocol-specific accuracy for selected anatomical and kinematic measurements [42]. The evidence map records which quantity was tested, its reference method, and the conditions under which that accuracy applies.

4 Computational biomechanics of the reconstructed wrist

Computational biomechanics estimates how a candidate reconstruction responds to specified tasks. In-vivo and experimental observations identify relevant outputs (Table 4); MBS and FEA relate patient geometry, retained tissues, loading, contact, and fixation under stated assumptions.

Tab. 4.  Wrist mechanical findings and supported comparisons

| Study | Evidence setting | Measured or predicted response | Comparison supported | Interpretive scope |
| --- | --- | --- | --- | --- |
| Akhbari et al.[14] | In-vivo biplane videoradiography | Task-based COR excursion after TWA and in healthy wrists | Compare motion using the same algorithm and tasks | Device-, task-, and algorithm-specific |
| Akhbari et al.[13] | In-vivo contact analysis | Contact offset and migration on radial and carpal surfaces | Assess contact path under the reported conditions | Observed in the studied implant and tasks |
| McCullough et al. {McCullough 2012} | Finite element comparison | Articular shape and tendon forces change contact and stability | Compare geometry under matched model assumptions | Model- and input-dependent prediction |
| Gislason et al. {Gislason 2017} | Universal 2 finite element study | Predicted component loading during the modelled task | Relate device response to the stated load path | Universal 2 construct and stated conditions |
| Bain et al. {Bain 2025} | Cadaveric hexapod study | Stiffness, carpal height, and center displacement after implantation | Evaluate geometry and soft-tissue consequences in that construct | Cadaveric construct under tested conditions |
| Norman-Puim et al. {Norman-Puim 2026} | Cadaveric articulation comparison | Muscle force and carpal loading under tested motions | Compare polarity under matched tasks and boundary conditions | Tested articulation and motions |

4.1 Mechanical evidence from implanted wrists

TWA consolidates part of the carpus and concentrates articulation within an artificial joint. The reconstructed response therefore depends on the implant, retained tissues, and fixation as a single load path [8]. A FEA study of a Universal 2 predicted dorsal-ulnar component loading during grip under its modelled conditions [16, 17]. The task and model assumptions determine how that result transfers to another device or patient.

In-vivo measurements demonstrate that bearing geometry and functional motion can differ. A Freedom prosthesis cohort exhibited approximately 17.1 mm of COR excursion across the studied tasks compared with 7.2 mm in healthy wrists, alongside contact migration on the radial and carpal surfaces [13, 14]. COR position, contact path, offset and carpal height affect tendon moment arms, fixation moments and clearance [13–15, 60]. Cadaveric and simulator studies associate component geometry and articulation polarity with changes in stiffness, muscle forces, and carpal-component loading under their tested conditions [61]. These study-specific quantities support matched comparisons; clinically meaningful targets require outcome-linked evidence.

Clinical failure patterns identify the mechanical responses that computational evaluation should prioritize. Reported osteolysis, loosening, instability, and impingement focus computational assessment on fixation, bearing contact, clearance, and soft-tissue balance [6, 10, 12]. The model represents the task and implant architecture relevant to the failure mechanism under study.

4.2 Multibody and finite element model architecture

MBS represents bones and implant components as bodies linked by articulations, contacts, ligaments, and musculotendinous units. Model resolution follows the intended output: simplified carpal segments support selected joint-level analyses, while intercarpal motion requires corresponding articulations and constraints [22, 30, 59]. Ligament attachments, slack lengths, stiffness, muscle paths, and moment arms define the soft-tissue contribution to motion and stability [33].

FEA represents deformable components, supporting bone, bearings, and fixation through meshes, material laws, and interface conditions [62]. Sequential MBS-to-FEA transfer suits analyses in which local deformation has little effect on system motion. When compliance changes contact or load redistribution, the model updates coupled quantities and reports the convergence criterion [59, 63]. Initial fixation and osseointegrated fixation are distinct interface states and require separate assumptions.

4.3 From motion to loading

Loading analysis begins with a defined task. Grip, resisted flexion, deviation, dart-throwing motion, and pinch impose different muscle and joint-loading demands [7, 30]. Inverse dynamics estimates net moments for prescribed motion, and a recruitment rule distributes the muscular contribution according to force capacity and moment arms. Static optimization may omit stabilizing co-contraction unless the model constrains or measures it [64].

Each load case states the task, external force magnitude, direction and application point, prescribed motion, constrained degrees of freedom, recruitment objective, and soft-tissue preload. A free-body diagram determines whether the load enters through hand contact, individual muscle forces, or a joint resultant; applying the same resultant twice violates force balance. MBS-to-FEA transfer requires forces and moments in the same coordinate system and about specified reference points. Comparisons hold the external task constant while allowing internal forces to respond to candidate geometry.

4.4 From loading to local mechanical response

The following reporting template makes the boundary conditions needed to interpret candidate comparisons explicit. It is a proposed extraction tool; study-specific values should be transcribed from the original methods.

Tab. 5.  Model conditions required for reproducible TWA comparisons

| Model field | Report | Why it changes interpretation |
| --- | --- | --- |
| Task and kinematics | Task, posture, motion prescribed or solved, and time or loading cycle | Defines the response being predicted |
| External loading | Magnitude, direction, application point, waveform, and source | Sets the resultant and task comparability |
| Muscles and ligaments | Included structures, recruitment rule, force-length or stiffness inputs, and preload | Changes joint forces and stability |
| Coordinate transfer | Anatomical frames, rotation convention, units, moment reference point | Prevents force and moment transfer errors |
| FEA boundary | Subdomain, truncated anatomy, fixed/loaded surfaces, constrained degrees of freedom | Defines support and prevents unintended rigid motion |
| Bearing contact | Contact pairs, separation, friction/sliding law, and solver assumptions | Controls pressure, migration, and edge response |
| Implant-bone interface | Initial fit, interface state, bonding/sliding law, and parameter source | Determines whether local micromotion can be predicted |
| Material and mesh | Constitutive law, patient-specific calibration, element type, mesh study | Controls stiffness and numerical solution |
| Validation | Independent reference, matched condition, output, and error measure | Separates input reproduction from predictive validation |

Transfer each MBS resultant into the FEA frame before applying it. Let R_BA rotate vector components from frame A to frame B; let r_(O_B to O_A) point from the new moment origin to the original origin.

F^B = R_BA F^A

M^B_O_B = R_BA M^A_O_A + r^B_(O_B to O_A) × F^B

State units and preserve the resultant force and moment when a distributed load is reduced to an equivalent load.

For a continuum model, state the balance law and whether the task is treated as quasi-static:

∇ · σ + b = ρ ü

Specify displacement and traction boundaries separately:

u = ū on Γ_u;  σ n = t̄ on Γ_t

Report the material law, loaded and fixed surfaces, and constrained degrees of freedom.

For frictional contact, define the normal gap g_n, compressive pressure p_n, and tangential traction t_t:

g_n ≥ 0;  p_n ≥ 0;  g_n p_n = 0;  ||t_t|| ≤ μ p_n

Local interface motion uses implant and bone displacements at matched locations:

Δu = u_implant − u_bone;  δ_n = Δu · n;  δ_t = ||Δu − δ_n n||

The interface location, reference frame, and loading interval define these outputs; state the stick-slip law used for tangential contact.

In an MBS-FEA workflow, task-specific forces, moments, and component poses define the FEA configuration and loads [17]. Contact area, pressure distribution, contact migration, edge clearance, component stress, bone strain, and fixation-interface motion describe distinct mechanical responses [13]. Each output is reported for a defined region, load case, and unit.

Interface micromotion is the local relative displacement of implant and bone at matched interface locations, resolved into normal and tangential components for a defined loading state. Component stress, screw force, and bone strain describe structural demand; contact pressure, sliding distance, and cross-shear support wear estimation when calibrated to the bearing material and loading history [65]. Postoperative migration is the change in component position relative to bone over time and is a separate measure.

4.5 Model verification validation and uncertainty

Numerical verification tests solution accuracy for the outputs used to rank designs. FEA studies report mesh and contact convergence for pressure, regional strain, and interface displacement, together with material assignment and cortical representation [59]. Coupled models also require consistent coordinates, load transfer, and force balance.

Validation compares each prediction with an independent measurement under matched conditions. Dynamic imaging and cadaveric tracking assess kinematics; instrumented loads, stiffness, or contact observables assess mechanical response [66, 67]. A prescribed pose is an input, so reproducing it does not independently validate pose prediction.

Uncertainty analysis propagates supported variation in geometry, placement, tissue properties, recruitment, and fixation into prediction intervals [3]. Sensitivity analysis identifies assumptions that change candidate rankings. The decision record reports the output, uncertainty, evidence basis, and intended use [18, 68].

4.6 Translating predictions into design decisions

Motion trajectories and functional arcs define comparison objectives [4]; contact, interface motion, component stress, and screw loading define mechanical responses [16]. Numerical values become acceptance criteria only when a condition-matched source supports that use. The design record states the quantity, interval, task, interface state, and evidence behind each objective.

Fig. 2. Mechanical quantities and corresponding assessment methods. MBS transfers task loads and poses to a defined FEA subdomain; model outputs are compared with matching construct or imaging measurements. Postoperative migration, task-based kinematics, and clinical outcomes occupy distinct timescales and require separate assessment.

5 Implant design optimization and customization

Design evaluation compares candidate geometries against patient-specific requirements and the mechanical responses defined in Section 4. Changes in bearing shape, fixation, or placement return to the same task and model conditions for comparison.

5.1 Defining the design space

Diagnosis, bone quality, activity, and prior surgery define the required balance among motion, stability, fixation, and durability [36]. Design variables include conformity, curvature, coverage, edge geometry, bearing position, footprint, stem and screw paths, porous regions, and resection volume [14]. Each variable maps to a clinical requirement and measurable response; constraints include tissue clearance, available bone, surgical access, and manufacturing tolerance.

5.2 Articular geometry fit and fixation

Articular geometry shapes motion, contact, and soft-tissue loading. Spherical, ellipsoidal, ovoid, and toroidal surfaces provide different curvature profiles [62]. Translation and COR behavior also depend on conformity, clearance, alignment, and tissue constraints. Candidate geometry is compared across defined tasks using trajectory, contact pressure, stability, impingement, and fixation moments. Centroid and orientation errors can enter an objective function after normalization by measurement uncertainty and clinical relevance [23].

External geometry should provide cortical support, suitable canal engagement, adequate thickness and clearance from vulnerable structures. Local underhang, overhang and fixation-corridor measurements make these requirements explicit [28, 56]. Articular position, fixation geometry and resection should be optimized together because they jointly determine load paths and retained bone. SSM-derived virtual cohorts support modular sizing and population-coverage evaluation, with the source population and sampling strategy recorded [55, 56].

5.3 Optimization durability and acceptance criteria

Multiobjective optimization compares motion error, contact loading, interface motion, bone resection, wear, and fixation demand [22, 59, 62]. Pareto solutions expose trade-offs; constraints define engineering requirements such as bearing thickness, fatigue performance, screw clearance, manufacturability, and surgical access. Surrogate models can reduce computational cost when their accuracy is verified over the candidate design space. Ligament retention, tendon tension, capsule closure and carpal height remain within the evaluation since they affect the assembled reconstruction [1, 66]. Robust comparisons vary anatomy, placement, and loading within supported ranges and report when candidate rankings change. Objective weights and clinical preferences should be stated explicitly.

Wear evaluation uses multidirectional motion and a defined loading history because polyethylene response depends on pressure, sliding, cross-shear, surface finish, and third-body effects [65]. An Archard-type formulation estimates wear when calibrated to the material and protocol. Design reports define how trajectory error, edge clearance, contact pressure, interface motion, fatigue, wear, and impingement are assessed [22, 30, 31, 59]. Values transferred from other joints remain tied to their source conditions [65, 69].

5.4 Customization design release and surgical planning

Customization spans size matching, anatomical customization and functional customization [13, 14]. Size matching selects a modular component; anatomical customization adapts external geometry and fixation; functional customization changes bearing shape, position or constraint to achieve a specified mechanical response. Fit and function require separate assessment because their objectives can compete. Fully individualized components are particularly relevant when modular systems cannot provide adequate fixation, alignment or bone preservation [39]. The implementation route should be planned alongside the design inputs for individualized devices [28].

Wei et al. reported contralateral-CT-based customized TWA in four severely damaged wrists, using metallic components, a polyethylene bearing and porous fixation [28]. A separate fourteen-wrist series reported improvements in pain, grip and motion at approximately twenty months (19.7±10.7 months), with no observed loosening during that follow-up [29]. These series establish short-term clinical feasibility within their reported cohorts [2, 8]. A quasi-dynamic FEA comparison, solving successive task postures, reported more uniform stress and steadier COR behavior for a customized design within its modelling conditions [62]. Custom scaphoid and lunate replacements provide related reconstruction experience, while their articulation and fixation differ from TWA [70, 71]. Independent testing and comparative follow-up connect these feasibility results to decisions about wider use.

Design release follows approval of reconstructed anatomy, virtual implantation, tolerance analysis, mechanical assessment and manufacturing review [28]. The radial component coordinates bearing position, canal engagement, cortical support and distal radioulnar clearance. The carpal component coordinates bone anchorage, screw trajectories, thickness, tendon clearance and revision options [28, 72]. Porous fixation regions and polished bearing surfaces require separate functional specifications [73]. The operative plan defines resection planes, tolerances, proximal-row treatment, retained tissues, screw placement and trial reduction with soft-tissue balancing [1, 32, 67]. Activity guidance should be consistent with the loading envelope used for design assessment [29, 36]. Source images, models, design versions and approvals form the release record.

6 Materials manufacturing and implementation

Stage 5 realizes the selected design in a device whose material properties, dimensions and surfaces can be inspected. Process qualification and the digital design record connect predicted performance to the manufactured implant.

6.1 Materials bearings and fixation surfaces

Material selection follows the requirements of the bearing, structural body, fixation interface and modular junctions [6, 62]. Cobalt-chromium-molybdenum offers hardness and finish retention for polished bearings, with fatigue and corrosion behavior influenced by composition and processing [74]. Its high elastic modulus motivates attention to stiffness and load transfer at fixation interfaces [73, 74]. Titanium alloys support structural and porous bone-contact regions [73, 75]; their direct bearing use requires additional tribological justification [73, 75, 76]. Polyethylene formulations balance wear resistance, oxidation resistance and fatigue performance under thin-section, multidirectional loading [76–78]. PEEK and reinforced PEEK offer alternative stiffness and imaging properties, with formulation-dependent anisotropy and debris behavior [7, 79]. Metal-on-metal designs require assessment of wear, corrosion and tissue exposure informed by relevant arthroplasty experience [69, 80]. Porous and calcium-phosphate-treated surfaces support biological fixation through architecture, surface properties and initial stability [81, 82]. Cemented and cementless fixation should be evaluated as construct-specific choices [10, 39].

6.2 Process qualification and dimensional inspection

Conventional manufacturing combines casting, forging or machining with appropriate heat treatment, finishing, cleaning and inspection [75, 73]. Bearing surfaces require control of form, waviness, roughness and edge transitions, with porous fixation features kept outside the articulation. Additive manufacturing enables individualized shapes and porous structures, while process qualification addresses powder history, anisotropy, residual stress, defects, surface condition and support removal [73, 74, 82, 83]. Calibrated dimensional and surface measurements should verify local curvature and the relative positions of bearing and fixation features. As-built porous architecture requires inspection because its stiffness and fatigue behavior follow the manufactured structure [74, 75, 82]. First-article and production inspection records complete the connection between the approved geometry and delivered component.

6.3 Digital traceability and implementation

The digital thread links source imaging, segmentation, reconstruction, design approval, manufacture, inspection and surgical planning [28, 55, 62]. Patient identity, side, coordinate systems and version status should remain consistent across these transformations. Human review of critical anatomy and fixation corridors, together with change control, keeps the produced component aligned with the approved plan [55]. Cleaning, sterilization and packaging require validation compatible with porous structures, polymers and surface treatments [75]. The applicable implementation route determines the required biocompatibility, risk-management and device evidence [29]. Materials and manufacturing experience from other orthopaedic devices provide a starting point for wrist-specific qualification. Component and construct testing then establish whether the manufactured implant meets its design requirements.

7 Design verification clinical validation and surveillance

Stage 6 combines testing of the manufactured reconstruction with assessment of delivered clinical function. Verification tests specified design requirements, while validation assesses whether the reconstruction meets its intended patient needs. Postoperative surveillance links these assessments to subsequent design and model updates.

7.1 Mechanical verification and testing

Testing progresses from material specimens to components, assembled constructs and cadaveric reconstruction [31, 82]. Material tests characterize strength, fatigue, corrosion and porous architecture; component tests evaluate cyclic loading, insert retention and fixation features [62]. Construct tests should reproduce the tasks of Section 4.3, including preload, cyclic histories and placement variation, with specimen characteristics and pass criteria defined in advance [16, 17]. Cadaveric simulators and imaging assess resulting kinematics and tendon-force requirements [66, 84]. Instrumented constructs provide controlled load measurements against which models can be assessed, while direct in-vivo load measurement remains an important development priority [7].

Wear testing should combine representative multidirectional motion with defined loading histories. Gravimetric or volumetric wear, surface damage and particle morphology characterize the response, with lubricant and temperature conditions recorded. Material-level tribometry complements construct testing [65]. The assembled reconstruction should also be evaluated for impingement involving stems, necks, screws and component edges [1]. Fixation tests should assess cyclic stability and interface response under the same task-based loading assumptions used in analysis.

7.2 Retrieval analysis

Retrieval analysis provides physical evidence of the loading and damage experienced in clinical use. Bearings, modular junctions and fixation features can be examined for wear, corrosion, fracture and interface changes [1]. Operative findings and surrounding tissue provide additional information on loosening, bone loss and tendon complications [85, 86]. Comparison with simulator damage tests has been used to propose damage mechanisms and refine loading protocols. Recording extraction performance and retained bone also informs future revision-oriented design [39, 40].

7.3 Postoperative measurement and longitudinal validation

Postoperative imaging measures alignment, tissue change, and component motion. Radiographs identify lucency and gross position. Wrist-specific photon-counting CT has been evaluated for component migration in cadaveric and clinical settings; RSA evidence from other arthroplasty sites can inform protocol development when transfer is tested for the wrist [87, 88]. Metal-artifact reduction is assessed for the selected implant and acquisition protocol [89]. Clinical interpretation pairs precision estimates with outcome-linked measures of loosening and patient function. Biplanar videoradiography measures task-based kinematics for comparison with predicted motion [90].

Ultrasound interface methods measure acoustic responses to controlled changes in bone-implant systems. Interferometric work estimated interlayer thickness in idealized systems and a bone-implant test setup [91]. Guided-wave experiments measured responses to simulated interface-stiffness changes in a synthetic model of a femoral transcutaneous implant [92]. These studies establish experimental methods for their test systems.

Clinical validation records pain, grip strength, motion, patient-reported function, complications, and durability. PRWE and MHQ provide standardized patient-reported measures [93, 94]. Reports state indication, device, follow-up, revision definition, and outcome timing so that comparisons retain their clinical context [11]. Paired clinical and mechanical measurements can test whether a prespecified response is associated with function or failure.

Longitudinal assessment pairs predicted wear, fixation behavior, and mechanical response with imaging, retrieval, and clinical observations [75]. The central hypothesis is that a reconstruction meeting defined functional and mechanical requirements will support durable patient benefit. Testing this relationship requires credible loading estimates, comparative evaluation of functional customization, wrist-relevant wear protocols, and prospective follow-up linking mechanical measures to patient outcomes [3, 7, 15]. Existing reports establish individual capabilities, while the relationship between measured mechanical fidelity and long-term survival remains a prospective research question [10, 13, 14].

8 Design decisions and research priorities

The six-stage synthesis organizes evidence by the decision each method can support. Table 6 distinguishes wrist measurements, wrist models, clinical reports, and methods transferred from other settings. The priority is to connect these evidence types through comparable tasks and defined outputs.

8.1 Evidence maturity and design decisions

Native mechanics and clinical observations define functional references and reconstruction requirements. Validated imaging and tracking protocols provide quantitative anatomy and motion measurements [20, 26], while segmentation methods support their preparation [42, 47, 49, 50]. Computational studies estimate selected loading and contact responses [22, 59, 63], and customized implant reports demonstrate feasibility [28, 62]. Manufacturing experience supplies qualification methods [74, 75, 82], and surveillance developments enable quantitative postoperative measurement [87, 88]. The principal translational task is to link these methods through consistent quantities, credible loading assumptions and prospective comparison of predicted and delivered performance.

Design decisions account for implant generation, indication, surgical practice, and follow-up when interpreting failure patterns [8, 10]. The proposed matrices make these conditions visible and provide a common record for future comparative workflows. Near-term progress depends on comparable task and boundary-condition reporting, independent assessment of model outputs, and prospective linkage of mechanical measurements with patient outcomes.

Tab. 6.  Evidence map by decision and applicability

| Evidence area | Representative sources | Supports | Evidence role | Applicability |
| --- | --- | --- | --- | --- |
| Clinical requirements | TWA reviews and cohorts[2, 8] | Indication, outcomes, complications, comparator context | Clinical evidence | Device- and population-specific |
| Native wrist mechanics | Anatomy, daily tasks, motion references[3, 5] | Functional references and task definitions | Healthy-wrist context | Reference for task definition |
| Imaging and tracking | CT, MRI, biplanar validation[8, 23, 42, 90] | Anatomy and task-based bone pose | Protocol-specific measurement accuracy | Accuracy applies to tested protocol |
| Segmentation | Shape models and learning methods[47, 49, 55] | Reconstructed surfaces and landmarks | Validation targets decision-critical anatomy | Report anatomical target and error measure |
| Wrist MBS | Wrist musculoskeletal models[22, 30, 59] | Conditional muscle and joint force estimates | Model-dependent estimates | Task and recruitment assumptions matter |
| TWA FEA | Implant and bearing studies[16, 17] | Contact, stress, load transfer | Device-specific predictions | Boundary conditions define comparison |
| In-vivo mechanics | TWA tracking/contact studies {Akhbari 2020; Akhbari 2021} | COR trajectories and contact migration | Observed measurements | Reported tasks and systems |
| Customization | Custom TWA cohorts and comparisons {Wei 2025; Cai 2023; Xiao 2026} | Design rationale, feasibility, modelled responses | Selected clinical reports | Comparative benefit remains an outcome question |
| Manufacturing | Materials and process studies {Wang 2016; Mani 2024; Kutluay 2026} | Process qualification and component properties | Process and component evidence | Requires device-specific qualification |
| Wrist migration imaging | Photon-counting CT/RSA wrist study {Lind 2026} | Component migration precision | Direct wrist measurement evidence | Clinical meaning requires outcome evaluation |
| Interface ultrasound | Interferometry and synthetic transcutaneous model {Lützelberger 2023; Chen 2026} | Acoustic response to controlled interface conditions | Methods from experimental systems | No wrist TWA validation reported |

8.2 Patient-specific design and longitudinal assessment

A longitudinal patient-specific model connects preoperative anatomy and targets with executed placement and postoperative observations. Candidate update quantities include task-based motion and COR behavior [4, 21], contact migration [13], component migration [87, 88], and clinical function [10, 36]. Imaging, segmentation, coordinate systems, design revisions and inspection records provide the digital thread supporting those updates [25, 28, 62]. Model credibility and virtual-cohort comparisons define the decisions the updated model can support [55, 68].

Customized TWA reports describe selected cases and model comparisons [25, 28, 62]. Comparative studies report the indication for customization, geometry changed, predicted response, operative feasibility, and postoperative outcome, using a relevant modular comparator when available [8, 18].

Population-based sizing and individual adaptation address complementary problems. Statistical shape models support component families and virtual-cohort assessment; functional adaptation changes curvature, bearing position, or constraint to target a defined task [55, 56, 62]. Manufacturing qualification and extraction planning connect these designs to reproducible components and future revision options [40, 82].

8.3 Research priorities

Research priorities are standardized definitions of motion, contact, and migration; benchmark data for segmentation and tracking; task-based reporting of loads and boundary conditions; and wrist-specific construct tests linked to stated design criteria [7, 26]. Prospective multicenter studies pair these mechanical measures with function and survival while recording diagnosis, bone quality, activity, and surgical experience [2, 8, 18]. These comparisons can determine which design choices improve patient-relevant outcomes.

9 Conclusions

Taken together, the evidence reviewed suggests that the principal limitation of total wrist arthroplasty is not hardware alone, but the absence of an integrated methodology. Successive implant generations have improved pain relief and preserved motion, yet have shifted rather than eliminated failure modes. Reported survivorship also continues to lag behind that of major joint arthroplasty. Because the wrist is a patient specific motion system, reconstructive success must be judged by more than component position or fit. It depends on restoring motion and load transfer in relation to individual anatomy and kinematics.

Improvement requires a traceable pathway that translates clinical needs into measurable requirements, incorporates imaging and motion data with their uncertainties, tests computational models against observed wrist behavior, and documents verification from design through clinical evaluation. The literature reviewed describes many of these capabilities, but does not yet demonstrate their consistent integration across this full sequence. Connecting them is therefore a central priority. Measurement and design should advance together, with critical attention to the handoffs between disciplines.

Total wrist arthroplasty is a stringent test of personalized implant medicine because clinical volumes are low, mechanics are complex, and the margin for error is narrow. A patient specific approach may also inform reconstruction of other small, complex joints. Its objective is to restore physiological motion and load, preserve bone and future revision options, and treat each implant as a mechanical hypothesis about an individual patient, tested through postoperative measurement. The field’s next advance depends on turning implant iteration into systematic learning.

References

1.	Reigstad O (2025) Total wrist arthroplasty: recent advances and current recommendations. J Hand Surg Eur Vol 50:34–41. https://doi.org/10.1177/17531934241276383

2.	Berber O, Garagnani L, Gidwani S (2018) Systematic Review of Total Wrist Arthroplasty and Arthrodesis in Wrist Arthritis. J Wrist Surg 7:424–440. https://doi.org/10.1055/s-0038-1646956

3.	Eschweiler J, Li J, Quack V et al. (2022) Anatomy, Biomechanics, and Loads of the Wrist Joint. Life (Basel) 12. https://doi.org/10.3390/life12020188

4.	Moore DC, Crisco JJ, Trafton TG et al. (2007) A digital database of wrist bone anatomy and carpal kinematics. J Biomech 40:2537–2542. https://doi.org/10.1016/j.jbiomech.2006.10.041

5.	Moritomo H, Apergis EP, Herzberg G et al. (2007) 2007 IFSSH committee report of wrist biomechanics committee: biomechanics of the so-called dart-throwing motion of the wrist. J Hand Surg Am 32:1447–1453. https://doi.org/10.1016/j.jhsa.2007.08.014

6.	Kennedy CD, Huang JI (2016) Prosthetic Design in Total Wrist Arthroplasty. Orthop Clin North Am 47:207–218. https://doi.org/10.1016/j.ocl.2015.08.018

7.	Eschweiler J, Li J, Quack V et al. (2022) Total Wrist Arthroplasty-A Systematic Review of the Outcome, and an Introduction of FreeMove-An Approach to Improve TWA. Life (Basel) 12. https://doi.org/10.3390/life12030411

8.	Zijlker HJA, Ritt MJPF, Beumer A (2022) Fourth-Generation Total Wrist Arthroplasty: A Systematic Review of Clinical Outcomes. J Wrist Surg 11:456–464. https://doi.org/10.1055/s-0041-1735840

9.	Krukhaug Y, Lie SA, Havelin LI et al. (2011) Results of 189 wrist replacements. A report from the Norwegian Arthroplasty Register. Acta Orthopaedica 82:405–409. https://doi.org/10.3109/17453674.2011.588858

10.	Newton A, Kandemir G, Joyce T et al. (2023) Long-term outcomes of the Universal 2 total wrist replacement: revision and loosening at 10 years and beyond. J Hand Surg Eur Vol 48:641–647. https://doi.org/10.1177/17531934231160380

11.	Taylor F, Gilpin BD, Sivakumar BS et al. (2024) Increasing Use of Total Wrist Arthroplasty-An Australian National Joint Registry Report. J Wrist Surg 13:339–345. https://doi.org/10.1055/s-0043-1777408

12.	Boeckstyns MEH, Herzberg G (2024) Complications after total wrist arthroplasty. J Hand Surg Eur Vol 49:177–187. https://doi.org/10.1177/17531934231203297

13.	Akhbari B, Morton AM, Shah KN et al. (2021) In vivo articular contact pattern of a total wrist arthroplasty design. J Biomech 121:110420. https://doi.org/10.1016/j.jbiomech.2021.110420

14.	Akhbari B, Morton AM, Shah KN et al. (2020) Proximal-distal shift of the center of rotation in a total wrist arthroplasty is more than twice of the healthy wrist. J Orthop Res 38:1575–1586. https://doi.org/10.1002/jor.24717

15.	Bain GI, Bellringer SF, Russo MP et al. (2025) Biomechanical analysis of cadaveric wrists before and after MOTEC wrist arthroplasty using a hexapod robot. J Hand Surg Eur Vol 50:492–499. https://doi.org/10.1177/17531934241274135

16.	Gislason MK, Foster E, Bransby-Zachary M et al. (2017) Biomechanical analysis of the Universal 2 implant in total wrist arthroplasty: a finite element study. Comput Methods Biomech Biomed Engin 20:1113–1121. https://doi.org/10.1080/10255842.2017.1336548

17.	McCullough MBA, Adams BD, Grosland NM (2012) THE EFFECT OF ARTICULAR SURFACE SHAPE AND TENDON FORCES OF TOTAL WRIST ARTHROPLASTY SYSTEMS: A FINITE ELEMENT STUDY. J Musculoskelet Res 15:1250021. https://doi.org/10.1142/S0218957712500212

18.	Pappalardo G, Schäfer L, Falco D et al. (2026) Total wrist arthroplasty efficacy and safety: a systematic review. J Orthop 77:49–58. https://doi.org/10.1016/j.jor.2026.04.001

19.	Smith AM, Maling LC, Williamson M (2024) Total wrist arthroplasty: commentary and opinions. J Hand Surg Eur Vol 49:115–118. https://doi.org/10.1177/17531934231209638

20.	Pfaeffle J, Blankenhorn B, Stabile K et al. (2005) Development and validation of a computed tomography-based methodology to measure carpal kinematics. J Biomech Eng 127:541–548. https://doi.org/10.1115/1.1894370

21.	Akhbari B, Moore DC, Laidlaw DH et al. (2019) Predicting Carpal Bone Kinematics Using an Expanded Digital Database of Wrist Carpal Bone Anatomy and Kinematics. J Orthop Res 37:2661–2670. https://doi.org/10.1002/jor.24435

22.	Eschweiler J, Stromps J-P, Fischer M et al. (2016) Development of a biomechanical model of the wrist joint for patient-specific model guided surgical therapy planning: Part 1. Proc Inst Mech Eng H 230:310–325. https://doi.org/10.1177/0954411916632791

23.	Li J, Rath B, Hildebrand F et al. (2022) Wrist Bone Motion during Flexion-Extension and Radial-Ulnar Deviation: An MRI Study. Life (Basel) 12. https://doi.org/10.3390/life12101458

24.	Jia S, Xiang Z, Lu X et al. (2026) Quantifying in vivo kinematics of the wrist during opposite dart-throwing motion: a 4D CT study. J Orthop Surg Res 21. https://doi.org/10.1186/s13018-026-06810-7

25.	Zhao K, Breighner R, Holmes D et al. (2015) A technique for quantifying wrist motion using four-dimensional computed tomography: approach and validation. J Biomech Eng 137:745011–745015. https://doi.org/10.1115/1.4030405

26.	Crisco JJ, Morton AM, Lopez C et al. (2026) Kinematic tracking of the small bones of the wrist in sequential 3DCT and dynamic 4DCT volume images using open-source Hierarchical 3D Registration, a module within SlicerAutoscoperM. BioMedical Engineering OnLine 25. https://doi.org/10.1186/s12938-026-01601-x

27.	Teule EHS, Lessmann N, van der Heijden EPA et al. (2024) Automatic segmentation and labelling of wrist bones in four-dimensional computed tomography datasets via deep learning. J Hand Surg Eur Vol 49:507–509. https://doi.org/10.1177/17531934231209876

28.	Wei Q, Liu C, Chen S et al. (2025) Customized Metal 3D Printed Total Wrist Prosthesis in the Treatment of Severely Destroyed Wrist: Design Rationale and Clinical Applications. Orthop Surg 17:1536–1546. https://doi.org/10.1111/os.70030

29.	Cai X, Wang B, Zhang B et al. (2022) Safety and efficacy of a novel three-dimensional printed microporous titanium prosthesis for total wrist arthroplasty in the treatment of end-stage wrist arthritis. Front Bioeng Biotechnol 10:1119720. https://doi.org/10.3389/fbioe.2022.1119720

30.	Eschweiler J, Praster M, Quack V et al. (2022) Musculoskeletal Modeling of the Wrist via a Multi Body Simulation. Life (Basel) 12. https://doi.org/10.3390/life12040581

31.	Mirakhorlo M, van Beek N, Wesseling M et al. (2018) A musculoskeletal model of the hand and wrist: model definition and evaluation. Comput Methods Biomech Biomed Engin 21:548–557. https://doi.org/10.1080/10255842.2018.1490952

32.	Hooke AW, Pettersson K, Sagerfors M et al. (2015) An anatomic and kinematic analysis of a new total wrist arthroplasty design. J Wrist Surg 4:121–127. https://doi.org/10.1055/s-0035-1549288

33.	Nichols JA, Bednar MS, Wohlman SJ et al. (2017) Connecting the wrist to the hand: A simulation study exploring changes in thumb-tip endpoint force following wrist surgery. J Biomech 58:97–104. https://doi.org/10.1016/j.jbiomech.2017.04.024

34.	Mercier J, Durdzinska Timoteo A, Baillot R et al. (2025) Distal Intersection Tenosynovitis: Surgical Insights From Five Cases. Journal of Clinical Medicine 14:2110. https://doi.org/10.3390/jcm14062110

35.	Boeckstyns MEH, Herzberg G, Merser S (2013) Favorable results after total wrist arthroplasty: 65 wrists in 60 patients followed for 5–9 years. Acta Orthopaedica 84:415–419. https://doi.org/10.3109/17453674.2013.823588

36.	Clementson M, Larsson S, Abramo A et al. (2024) Clinical and Patient-Reported Outcomes After Total Wrist Arthroplasty and Total Wrist Fusion: A Prospective Cohort Study with 2-Year Follow-up. JB JS Open Access 9. https://doi.org/10.2106/JBJS.OA.23.00081

37.	Gvozdenovic R, Boeckstyns M, Merser S (2020) Ulnar Head or Total Distal Radioulnar Joint Replacement, Isolated and Combined with Total Wrist Arthroplasty: Midterm Results. J Wrist Surg 9:411–416. https://doi.org/10.1055/s-0040-1712981

38.	Amundsen A, Rizzo M, Berger RA et al. (2023) Twenty-Year Experience With Primary Distal Radioulnar Joint Arthroplasty From a Single Institution. J Hand Surg Am 48:53–67. https://doi.org/10.1016/j.jhsa.2022.02.014

39.	Cash T, Talwalkar S (2025) Revision wrist arthroplasty strategies. J Hand Surg Eur Vol 50:42–49. https://doi.org/10.1177/17531934241297843

40.	Berber O, Gidwani S, Garagnani L et al. (2020) Salvage of the Failed Total Wrist Arthroplasty: A Systematic Review. J Wrist Surg 9:446–456. https://doi.org/10.1055/s-0040-1713728

41.	Ryu JY, Cooney WP, Askew LJ et al. (1991) Functional ranges of motion of the wrist joint. J Hand Surg Am 16:409–419. https://doi.org/10.1016/0363-5023(91)90006-W

42.	Li J, Nebelung S, Schock J et al. (2021) A Novel Combined Level Set Model for Carpus Segmentation from Magnetic Resonance Images with Prior Knowledge aligned in Polar Coordinate System. Comput Methods Programs Biomed 208:106245. https://doi.org/10.1016/j.cmpb.2021.106245

43.	Li J, Nebelung S, Rath B et al. (2020) A novel combined level set model for automatic MR image segmentation. Current Directions in Biomedical Engineering 6:20–23. https://doi.org/10.1515/cdbme-2020-3006

44.	Zarenia M, Arpinar VE, Nencka AS et al. (2022) Dynamic tracking of scaphoid, lunate, and capitate carpal bones using four-dimensional MRI. PLoS One 17:e0269336. https://doi.org/10.1371/journal.pone.0269336

45.	Naredo E, D'Agostino MA, Wakefield RJ et al. (2013) Reliability of a consensus-based ultrasound score for tenosynovitis in rheumatoid arthritis. Annals of the Rheumatic Diseases 72:1328–1334. https://doi.org/10.1136/annrheumdis-2012-202092

46.	Dobbe JGG, Roo MGA de, Visschers JC et al. (2019) Evaluation of a Quantitative Method for Carpal Motion Analysis Using Clinical 3-D and 4-D CT Protocols. IEEE Trans Med Imaging 38:1048–1057. https://doi.org/10.1109/TMI.2018.2877503

47.	Anas EMA, Rasoulian A, Seitel A et al. (2016) Automatic Segmentation of Wrist Bones in CT Using a Statistical Wrist Shape + Pose Model. IEEE Trans Med Imaging 35:1789–1801. https://doi.org/10.1109/TMI.2016.2529500

48.	Raith S, Deitermann M, Pankert T et al. (2025) Multi-label segmentation of carpal bones in MRI using expansion transfer learning. Phys Med Biol 70. https://doi.org/10.1088/1361-6560/adabae

49.	Isensee F, Jaeger PF, Kohl SAA et al. (2021) nnU-Net: a self-configuring method for deep learning-based biomedical image segmentation. Nature Methods 18:203–211. https://doi.org/10.1038/s41592-020-01008-z

50.	Ma J, He Y, Li F et al. (2024) Segment anything in medical images. Nature Communications 15:654. https://doi.org/10.1038/s41467-024-44824-z

51.	Ji C, Li J, Praster M et al. (2022) Smoothing the Undersampled Carpal Bone Model with Small Volume and Large Curvature: A Feasibility Study. Life (Basel) 12. https://doi.org/10.3390/life12050770

52.	Falcinelli C, Valente F, Vasta M et al. (2023) Finite element analysis in implant dentistry: State of the art and future directions. Dent Mater 39:539–556. https://doi.org/10.1016/j.dental.2023.04.002

53.	Brui E, Efimtcev AY, Fokin VA et al. (2020) Deep learning-based fully automatic segmentation of wrist cartilage in MR images. NMR Biomed 33:e4320. https://doi.org/10.1002/nbm.4320

54.	van de Giessen M, Foumani M, Streekstra GJ et al. (2010) Statistical descriptions of scaphoid and lunate bone shapes. J Biomech 43:1463–1469. https://doi.org/10.1016/j.jbiomech.2010.02.006

55.	Sarkalkan N, Weinans H, Zadpoor AA (2014) Statistical shape and appearance models of bones. Bone 60:129–140. https://doi.org/10.1016/j.bone.2013.12.006

56.	Cai X, Wu Y, Huang J et al. (2024) Application of statistical shape models in orthopedics: a narrative review. Intelligent Medicine 4:249–255. https://doi.org/10.1016/j.imed.2024.05.001

57.	Trentadue TP, Thoreson A, Lopez C et al. (2024) Morphology of the scaphotrapeziotrapezoid joint: A multi-domain statistical shape modeling approach. J Orthop Res 42:2562–2574. https://doi.org/10.1002/jor.25918

58.	Chen X, Graham J, Hutchinson C et al. (2013) Automatic inference and measurement of 3D carpal bone kinematics from single view fluoroscopic sequences. IEEE Trans Med Imaging 32:317–328. https://doi.org/10.1109/TMI.2012.2226740

59.	Eschweiler J, Stromps J-P, Fischer M et al. (2016) A biomechanical model of the wrist joint for patient-specific model guided surgical therapy: Part 2. Proc Inst Mech Eng H 230:326–334. https://doi.org/10.1177/0954411916635443

60.	Akhbari B, Shah KN, Morton AM et al. (2021) Total Wrist Arthroplasty Alignment and Its Potential Association with Clinical Outcomes. J Wrist Surg 10:308–315. https://doi.org/10.1055/s-0041-1725172

61.	Norman-Puim E, Lalone E, Kadar A et al. (2026) Reverse Total Wrist Arthroplasty: New Design and Biomechanical Advantages. J Hand Surg Am online ahead of print. https://doi.org/10.1016/j.jhsa.2026.06.005

62.	Xiao Y, Tan X, Liu J et al. (2026) Biomechanical properties of a new 3D-printed customized anatomical total wrist prosthesis : a quasi-dynamic finite element analysis. Bone Joint Res 15:311–321. https://doi.org/10.1302/2046-3758.153.BJR-2024-0425.R3

63.	Bavil AY, Karimi Dastgerdi A, Barrett RS et al. (2026) Evidence-informed methodological approaches for coupled neuromusculoskeletal-finite element analysis workflows in orthopaedic biomechanics. Front Bioeng Biotechnol 14:1875782. https://doi.org/10.3389/fbioe.2026.1875782

64.	Eschweiler J, Praster M, Quack V et al. (2022) Comparison of Optimization Strategies for Musculoskeletal Modeling of the Wrist for Therapy Planning in Case of Total Wrist Arthroplasty. Life (Basel) 12. https://doi.org/10.3390/life12040527

65.	Baykal D, Siskey RS, Haider H et al. (2014) Advances in tribological testing of artificial joint biomaterials using multidirectional pin-on-disk testers. Journal of the Mechanical Behavior of Biomedical Materials 31:117–134. https://doi.org/10.1016/j.jmbbm.2013.05.020

66.	Glanville J, Bates KT, Silver N et al. (2025) Replication of Coupled Movements of the Wrist: A Cadaveric Study of Total Wrist Arthroplasty. J Hand Surg Am 50:424–432. https://doi.org/10.1016/j.jhsa.2024.12.006

67.	Glanville J, Bates KT, Fichera S et al. (2026) Determining the ideal position of the longitudinal centre of rotation for a Motec® total wrist arthroplasty. J Hand Surg Eur Vol 51:179–187. https://doi.org/10.1177/17531934251378987

68.	Pappalardo F, Russo G, Tshinanu FM et al. (2019) In silico clinical trials: concepts and early adoptions. Briefings in Bioinformatics 20:1699–1708. https://doi.org/10.1093/bib/bby043

69.	Hall DJ, Pourzal R, Jacobs JJ (2020) What Surgeons Need to Know About Adverse Local Tissue Reaction in Total Hip Arthroplasty. J Arthroplasty 35:S55-S59. https://doi.org/10.1016/j.arth.2020.01.016

70.	Cioffi A, Rovere G, Bosco F et al. (2023) Treatment of Scaphoid Non-Unions with Custom-Made 3D-Printed Titanium Partial and Total Scaphoid Prostheses and Scaphoid Interosseous Ligament Reconstruction. Healthcare (Basel) 11. https://doi.org/10.3390/healthcare11243123

71.	Zhang C, Chen H, Fan H et al. (2023) Carpal bone replacement using personalized 3D printed tantalum prosthesis. Front Bioeng Biotechnol 11:1234052. https://doi.org/10.3389/fbioe.2023.1234052

72.	Reigstad O, Holm-Glad T, Korslund J et al. (2024) Long-Term Outcomes of Wrist Arthroplasty Using the ReMotion™ Implant in Non-inflammatory Wrist Pathology. J Hand Surg Asian Pac Vol 29:200–210. https://doi.org/10.1142/S2424835524500218

73.	Abd-Elaziem W, Darwish MA, Hamada A et al. (2024) Titanium-Based alloys and composites for orthopedic implants Applications: A comprehensive review. Materials & Design 241:112850. https://doi.org/10.1016/j.matdes.2024.112850

74.	Mani G, Porter D, Collins S et al. (2024) A review on manufacturing processes of cobalt-chromium alloy implants and its impact on corrosion resistance and biocompatibility. J Biomed Mater Res B Appl Biomater 112:e35431. https://doi.org/10.1002/jbm.b.35431

75.	Wang X, Xu S, Zhou S et al. (2016) Topological design and additive manufacturing of porous metals for bone scaffolds and orthopaedic implants: A review. Biomaterials 83:127–141. https://doi.org/10.1016/j.biomaterials.2016.01.012

76.	Wang A, Essner A, Polineni VK et al. (1998) Lubrication and wear of ultra-high molecular weight polyethylene in total joint replacements. Tribology International 31:17–33. https://doi.org/10.1016/S0301-679X(98)00005-X

77.	Spece H, Yarbrough RV, Kurtz SM (2023) In Vivo Performance of Vitamin E Stabilized Polyethylene Implants for Total Hip Arthroplasty: A Review. J Arthroplasty 38:970–979. https://doi.org/10.1016/j.arth.2022.11.010

78.	Bracco P, Oral E (2011) Vitamin E-stabilized UHMWPE for total joint implants: a review. Clin Orthop Relat Res 469:2286–2293. https://doi.org/10.1007/s11999-010-1717-6

79.	Huang H, Liu X, Wang J et al. (2024) Strategies to improve the performance of polyetheretherketone (PEEK) as orthopedic implants: from surface modification to addition of bioactive materials. J Mater Chem B 12:4533–4552. https://doi.org/10.1039/d3tb02740f

80.	Kwon Y-M, Della Valle CJ, Lombardi AV et al. (2021) Risk Stratification Algorithm for Management of Head-Neck Taper Tribocorrosion in Patients with Metal-on-Polyethylene Total Hip Arthroplasty: Consensus Statement of the American Association of Hip and Knee Surgeons, the American Academy of Orthopaedic Surgeons, and The Hip Society. J Bone Joint Surg Am 103:e18. https://doi.org/10.2106/JBJS.20.01837

81.	Karageorgiou V, Kaplan D (2005) Porosity of 3D biomaterial scaffolds and osteogenesis. Biomaterials 26:5474–5491. https://doi.org/10.1016/j.biomaterials.2005.02.002

82.	Kutluay GE, Erdoğan F, Dinçel YM (2026) Porous Ti-6Al-4V Architectures in Load-Bearing Orthopedic Reconstruction: A Critical Narrative Review of the Translational Gap Across Microstructure, Fatigue, Surface Function, and Clinical Failure Modes. Materials (Basel) 19:3860. https://doi.org/10.3390/ma19183860

83.	Li J, Shi D, Zhang Z et al. (2026) Data-driven meets knowledge-driven: an LLM‑agent framework for quality control in metal additive manufacturing. Journal of Industrial Information Integration 53:101182. https://doi.org/10.1016/j.jii.2026.101182

84.	Glanville J, Bates KT, Brown D et al. (2024) Evaluation of a cadaveric wrist motion simulator using marker-based X-ray reconstruction of moving morphology. PeerJ 12:e17179. https://doi.org/10.7717/peerj.17179

85.	Ward CM, Kuhl T, Adams BD (2011) Five to ten-year outcomes of the Universal total wrist arthroplasty in patients with rheumatoid arthritis. J Bone Joint Surg Am 93:914–919. https://doi.org/10.2106/JBJS.H.01614

86.	Wagner ER, Srnec JJ, Mehrotra K et al. (2017) What Are the Risk Factors and Complications Associated With Intraoperative and Postoperative Fractures in Total Wrist Arthroplasty? Clin Orthop Relat Res 475:2694–2700. https://doi.org/10.1007/s11999-017-5442-2

87.	Lind G, Tesselaar E, Sandberg O et al. (2026) Precision of photon-counting detector CT compared to conventional CT for total wrist arthroplasty migration: a cadaveric and clinical study. Skeletal Radiol 55:2835–2844. https://doi.org/10.1007/s00256-026-05293-1

88.	Jürgens-Lahnstein JH, Teilmann JF, Petersen ET et al. (2026) Inducible displacement measured with radiostereometric analysis in primary knee arthroplasty: a systematic review of clinical studies. Acta Orthopaedica 97:698–707. https://doi.org/10.2340/17453674.2026.46318

89.	Koo B, Kim D, Jung H et al. (2026) Quantitative phantom-based comparison of commercial CT metal artifact reduction algorithms. Journal of Applied Clinical Medical Physics 27:e70631. https://doi.org/10.1002/acm2.70631

90.	Akhbari B, Morton AM, Moore DC et al. (2019) Kinematic Accuracy in Tracking Total Wrist Arthroplasty With Biplane Videoradiography Using a Computed Tomography-Generated Model. J Biomech Eng 141:445031–445037. https://doi.org/10.1115/1.4042769

91.	Lützelberger J, Arneth P, Franck A et al. (2023) Ultrasonic Interferometric Procedure for Quantifying the Bone-Implant Interface. Sensors (Basel) 23:5942. https://doi.org/10.3390/s23135942

92.	Chen E, Fromme P (2026) Guided ultrasonic wave monitoring for osseointegration assessment of an intraosseous transcutaneous amputation prosthesis. Ultrasonics 163:107983. https://doi.org/10.1016/j.ultras.2026.107983

93.	Chung KC, Pillsbury MS, Walters MR et al. (1998) Reliability and validity testing of the Michigan Hand Outcomes Questionnaire. J Hand Surg Am 23:575–587. https://doi.org/10.1016/S0363-5023(98)80042-7

94.	MacDermid JC, Turgeon T, Richards RS et al. (1998) Patient rating of wrist pain and disability: a reliable and valid measurement tool. Journal of Orthopaedic Trauma 12:577–586. https://doi.org/10.1097/00005131-199811000-00009